## Slide 1: Implementation of Micram DAC4 AWG for Twin Pulse Output
Dr. Robert Sheehan, Chief Technical Officer, School of Physics
27 – 6 – 2025

## Slide 2: Twin Pulse Implementation
I was asked to configure the Micram DAC4 AWG to output the following pulse pattern in which Tw ≈ 100 (ps) and Td ≥ 100 (ps)
Pulse height V ≈ 400 mV, but this can always be amplified
Possible to obtain pattern with Tw ≈ 110 (ps) and Td ≥ 110 (ps)
Tw
Tw
Td

## Slide 3: Twin Pulse Implementation
An instability in the AWG output arises for clock input between 25 (GHz) and 29 (GHz)
Very high jitter observed in the output for clock input between 25 (GHz) and 29 (GHz), only one single pulse output visible also
I don’t know the cause, and don’t have time to investigate
System seems to re-stabilise for Fclck ≥ 29 (GHz)
Tw
Tw
Td

## Slide 4: Twin Pulse Implementation
Twin pulse can be implemented using Micram MATLAB code
Idea is that you specify the number of symbols in the pulse Nw≥ 5
Specify that number of symbols that make the delay Nd≥ 5
No output possible for Nw < 5
Use MATLAB to compute the number of symbols in the buffer Nb = 256 – 2 Nw – Nd
Signal Generator Input (Clock) specifies the actual output frequency
There’s an odd relationship between Clock Input and Operating output, something like Fout ≈ 0.133 Fclock

## Slide 5: Twin Pulse Implementation
Tw
Tw
Td
Tw controlled by Fclck and Nw
Td specified by choice of Nd
Measured Pulse Width
Measured Pulse Separation

## Slide 6: Twin Pulse Implementation
Tw
Tw
Td
Tw controlled by Fclck and Nw
Twmin ≥ 109.3 ± 8.3 (ps)
Measured Pulse Width
Measured Pulse Width (Zoomed In)

## Slide 7: Square Pulse Implementation
I implemented a square wave pulse output to examine the instability observed between 25 (GHz) and 29 (GHz)
The instability persisted in the same frequency range
I was able to show that it is possible to obtain pulses with Tw < 100 (ps)
Measurements show that square wave pulse widths are equal to the twin pulse widths, which is what you’d expect
Tw
Tw
Td
Tw
Tw
Td
Td

## Slide 8: Square Pulse Implementation
Tw
Tw
Td
Fout = 0.133 Fclck
Twmin ≥ 71.3 ± 4.2 (ps)
Tw
Tw
Td
Td
Measured Sq Wave Pulse Times
Relationship between Sq Wave Freq and Clock Input

## Slide 9: Square Pulse Implementation
Tw
Tw
Td
Td / Tw = 1.992 – 0.004 Fclck
Twmin ≥ 71.3 ± 4.2 (ps)
Tw
Tw
Td
Td
Measured Sq Wave Pulse Times
Data showing that Td = 2 Tw as required

## Slide 10: Conclusions
Using Micram DAC4 AWG it’s possible to obtain the desired twin-pulse output
It’s possible to control the pulse width and the pulse separation
Pulse width is controlled by the input clock
Output is stable for all frequencies outside the range 25 (GHz) to 29 (GHz)
Pulse widths down to 110 (ps) and less than 100 (ps) possible.
Timing measurements are limited by the resolution of the real-time scope ≈ 10 (ps)